## 2014 Budget
INEE Lab Budget 2014 (24 Dec 2013) Meeting
|  |  |  | Item | Detail Cost | Consumable | Service | VC's budget | Grant | Remark
| No | ID | Tool
| 1 | ICP RIE | Reactive ION Etching | Finger | 8000
oring
Sulfur hexafluoride
Carbon tetrafluoride
Oxygen
Trifluoromethane
Purified Argon
Helium
fomblin oil
Nitrogen Purge
purified air
|  |  |  | ceramic holder | 4000
| 2 | RTA | Repeat Thermal Anneal/RTO Upgrades. | MFC manual display and ball gas gauge | 900 |  |  |  | 500
|  |  |  | O2 Gas 5N | 1200 |  |  |  | 1200 | Ramzan's Grant
JT Lesker oil
O Ring
N2 Gas 5N
|  |  |  | 20 SLM Pump Edwards | 12000
| 3 | LPCVD | Low Pressure Chemical Vapour Dep | New Boat | 4000 |  |  | 4000
O Ring
Silane
Nitrogen Purge
Nitrogen Vent
fomlin oil
1900
|  |  |  | Scrubber | 6000 |  | 6000
|  |  |  | Rotary Pump Edward 80 | 2000 |  | 2000 |  |  |  | 2015 RM18000
| 4 | HT Eva | Thermal Evaporator | Kit Recycle | 300 x 3 |  |  | 900 |  | 3 sets
|  |  |  | O Ring | 900
|  |  |  | JT Lesker oil | 1.60 x 200 |  |  | 320
c |  |  |  | See thru glass | 4 x 900 |  | 3600 |  |  | 4 sets
|  |  |  | boat | 10=3100
|  |  |  | Chamber Sheild Recycle |  |  |  | 2000
| 5 | HP | Hot Plate | Sample Plate Holder upgrade
| 6 | SC | Spin Coater 1/2/3 | calibration | 3000
|  |  |  | O Ring | 500
| 7 | SP | Spinner
| 8 | USC | Ultra sonic Cleaner
| 9 | LT F | LT Furnace | glas tube | 3000
| 10 | Vac Ov | Vacuum Oven | Vacuum Pump | 2000
| 11 | PPS | Plasma Preen System | gas(N2 @ argon) | 600
|  |  |  | JT Lesker oil | 300
| 12 | MA | Mask Aligner | Mercury Lamp | 2500 |  |  | 2500
| 13 | Ox F | Wet/Dry Oxidation Furnace | gas (O2) | 200x2
|  |  |  | gas (O2) | 200x2
|  |  |  | gas (N2) | 200x2
|  |  |  | gas (N2) | 200x2
| 14 | MUF F | Muffle Furnace
| 15 | PS | DC Power Supply
| 16 | A Com | Air Compressor
| 17 | Si Com | Silent/Oil Compressor | tubing | 100
| 18 | Con Ov | Gravity/Natural Convection Oven
| 19 | Cent | 500g Centrifuge
| 20 | VO Mix | Vortex Mixer
| 21 | Bal | Balancing
| 22 | USC | ultra sonic cleaner DC200H
| 23 | A Clav | AutoClave | [7 empty] | 2015 RM1500
| 24 | Cent | Centrifuge 5430/5430 R
| 25 | PL | Photo Luminecence | service maintanance JKKP Cert
| 26 | SEM | JEOL SEM-EDS 6010LV | Contract | 24000 |  |  | 24000
|  |  |  | Filament | 6000 |  |  | 6000
| No | ID | Analysis
| 1 | Four Point Probe | Four Point Probe
| 2 | Osc | PC Oscilloscope | To be obsoleted.
| 3 | Osc | Oscilloscope | To be obsoleted.
| 4 | Hi-Pow | High Power Microscope 100X | service maintanance calibration
| 5 | Lo-Pow | Low Power Microscope 4.5X | To be obsoleted.
| 6 | 3D Profiler | 3D Surface Profilometer | calibration
| 7 | Die-analyzer | Dielectric Analyzer
| 8 | Sig-Gen | Signal Generator
| 9 | Elli | Ellipsometer | To be obsoleted.
| 10 | UV-IS | UV-IS Perkin Elmer UV/VIS Lambda 35
| 11 | PicoAmp | PicoAmpere
| 12 | Filmetric | Filmetric
| 13 | GAMRY | Potentiostat
| 14 | LCR | LCR Meter
| No | ID | Facility
| 1 |  | White Clean room booth | filter
| 2 |  | Yellow Clean room booth | filter
| 3 |  | DI Water | filter
| 4 |  | Chemical Bath
| 5 |  | Fume Hood in yellow booth
| 6 |  | Electronic Desiccators
| 7 |  | Micro-pipette
| 8 |  | Scrubber System | filter | 1500 x2
| 9 |  | chemical storage
| 10 |  | fume hood in nano material area
| 11 |  | Low Temprature Freezer
| 12 |  | Clarke
| 13 |  | Freezer
| 14 |  | DNA Laminar Flow Cabinet
| 15 |  | Gloves Box
| 16 |  | Water Pump Filter (building)
| No | ID | Supporting
| 1 |  | Chemical Handling
| 2 |  | Chemical Register | Positve Resist | 2 x 6000 | 12000 |  |  |  | Emi
|  |  |  | Negative R | 2 x 6000 | 12000 |  |  |  | Emi
|  |  |  | RD 6 Developer | 3000x4 | 12000 |  |  |  | Emi
|  |  |  | Acetone | 48 x 10 | 480 |  |  |  | Emi
|  |  |  | Buffered Oxide Etch (BOE) | 280x5 | 1400 |  |  |  | Che Omar
|  |  |  | Methanol | 200x5 | 1000 |  |  |  | Che Omar
|  |  |  | Ethanol | 100x6 | 600 |  |  |  | Che Omar
|  |  |  | Isopropanol (IPA) | 150 x 5 | 750 |  |  |  | Che Omar
|  |  |  | Aluminium Etch | 420x2 | 840 |  |  |  | Emi
|  |  |  | Hydrogen Peroxide | 250x2 | 500 |  |  |  | Emi
|  |  |  | Amonium Peroxide | 150x2 | 300 |  |  |  | Emi
|  |  |  | SU8 2005 | 4700x1 | 4700 |  |  |  | Emi
|  |  |  | SU8 Developer | 2200x1 | 2200 |  |  |  | Emi
|  |  |  | PDMS - Dow Corning Sylgard 184 | 2800x1 | 2800 |  |  |  | Emi
|  |  |  | HCL | 200x1 | 200 |  |  |  | Emi
| 3 |  | Gas Room | Aircondition | [5 empty] | Che Omar
| 4 |  | Chemical Cabinet |  |  | 51770
[5 empty] | Total Cost (RM) |  |  | 39720 | 1700
Note on G17: mira: Hook up completed on the 20th July 2012
Note on G23: mira: 18 May 2012 since April 2011.
Note on G76: mira: UP ( 10 aug 2012 )
Note on G82: mira: Comissioned on the 25th Feb 2013

## Lab Updates 2014
7 Feb 2014 |  | Kill the Snake Meeting (Last Updated)
| DATE | EQUIPMENTS/FACILITIES | ERROR CODE | PROBLEM DESCRIPTIONS | COUNTERMEASURE | ACTION/LEAD TIME | STATUS | PIC
1 | 14/6/2011 | LPCVD | -Chiller damaged | Chiller lost its pump and blower-Rundown condition | To order new or used chiller that is better and compatible | Unknown-Pending vendor to source out. TBD by WW02. Pending MARDI Blanket PO. WW14 All PO is out. 5-6 May carried out. Chiller replaced on the 6th May. | CLOSED | JASNI
2 | 14/6/2011 | LPCVD | -Tube over run | thick bi-product turned to charcoal | To recycle the parts at Kulim | Unknown-Pending vendor to confirm the LPCVD up and running.Pending MARDI Blanket PO.WW14 All PO is out. 5-6 May carried out. | CLOSED | JASNI
[5 empty] | WW04 Ulvac sourcing the chiller. Up on 18th June 2012 since April 2011. | Quotation out 3 Feb 2012. WW14 All PO is out. WW 12 WIP
[5 empty] | Pump E80 kaputt due to not flowing N2 dilution. Design problem due to N2 dilute didn't run automatically into the E80 pump. | Sending E80 for servicing, 2 wks leadtime, ETA 21 July. Completed on 21st July. | CLOSED | JASNI
[6 empty] | Finally up on the 30 th Aug, down again on the 3rd Sept due to pump burning. Re-coil the pump and pumping ok on the 22 Sept 2012 | CLOSED | JASNI
[6 empty] | Refurbish pump 16 Sept due to coil burnt | CLOSED | JASNI
3 | 2011-12-08 | 3D Profiler | -Software | Thickness cannot measure | Vendor come to repair but still failed. | TBD-Korean is doing the proposal | CLOSED | MIRA
[5 empty] | WW04 Magna Value Visit with Korean vendor doing the t/shooting. Hasrul follow up. | Problem solved due to software problem, but having hang problem. Limited to no reflective material. |  | MIRA
[6 empty] | Problem re-occured again. To reinstall program. 16 July, Some program left out in the software. Problem solved on the 8th August 2012. | CLOSED | MIRA
Hang again 7 Sept
4 | 22/8/2011 | Chemical Drainage system | -Leaking | No permanent fixing | Temporary, sub drainage. Permanent, ULVAC to engage or to follow up with Emi's counter measure with certified chemical vendor | Pending vendor to repairing the Chemical Drainage system | CLOSED | EMI
[5 empty] | Pending quotation from vendor | Install multiple filter.
5 | 2012-01-06 | Evaporator | Leak | Chamber glass chipping and crack | To order new chamber system | ULVAC is doing the costing & fabricating the new chamber. Pending MARDI Blanket PO. 5-6 May carried out. WW18 WIP. Closed 3rd June 2012. | CLOSED | JASNI
| 2012-01-21 |  |  |  | WW04 Ulvac sourcing the new vacuum chamber. | Glass chamber taken by ULVAC for vacuum chamber fabrications.
6 | 2012-01-06 | Evaporator/ LPCVD | Process Variable | No standard Process monitoring | To establish standard format | Require thickness measurement tool . Limitations to Thickness above 5Ams. | PENDING | MIRA
7 | 2012-01-06 | Tool Box | No right Tools for maintenance | Not Available | To buy 1 set | Quotation ready, pending Prof Uda to approve. Pending MARDI Blanket PO. WW14 All PO is out. | CLOSED | HASRUL
8 | 2012-01-30 | compressor | No proper recording system | cetificate N/A | To seek advise from UKKP Ishak Jainoo | Referring to DOSH but no clear indication. WW14 Pending Meeting with UKKP updates. | PENDING | HASRUL
9 | 2012-01-30 | General Exhaust/Scrubber | Not certified by JAS | cetificate N/A | To submit the plan JAS | Unknown | PENDING | ISA
10 | 2012-01-30 | Chemical Appron | Safety Hazard |  |  | Pending Quotation WW06 Magna Value. WW17 To buy. | CLOSED | HASRUL
11 | 2012-01-30 | Chemical boot (Welley's) | Safety Hazard |  |  | Pending Quotation WW06 Magna Value. WW17 To buy. | CLOSED | HASRUL
12 | 2012-01-30 | Sandals | Dead circuit | Shortage due to increasing number of Lab user | To be a dozen | Pending Quotation WW06 Magna Value. WW17 To buy. | CLOSED | HASRUL/ISA
13 | 2012-02-07 | Exhaust Furnace | Leak | detach from wall | Temporary rap with Al foil | Pending Vendor to fix WW06. | OPEN | JASNI
14 | 2012-02-07 | Ellipsometer | To bring over from MikroE Lab | Machine down since 2006 | To comm with SIRIM Vendor | Abort comm due to SIRIM also has no technical support. Move in WW06.3 | CLOSED | JASNI
Sourcing for Solution. Few questionable problems under considerations.
WW08 Updates, CC Kong from Functionstech is checking if the Laser Detector is the Problem.
| 2012-03-03 |  |  |  |  | Engaged KLA Tencor to assist investigations, taken away on WW10.5
WW14 Pending importing
| 2012-03-27 |  |  |  |  | 2nd measurement okay, pending confirmation. parts from US. Returned to INEE on 18th June, but still measurement.
| 2012-07-16 |  |  |  |  | Suspect Windows corrupt, to reinstall.
| 2012-11-26 |  |  |  |  | To dispose from INEE Property |  | JASNI
15 | 2012-02-08 | DI Water | Filter damaged | water cannot flow | To comm with Adscitech Vendor | Unknown. Move in WW06.1. WW10 Change the filter. | CLOSED | HASRUL
16 | 2012-02-15 | Power Trip Booth E-F | Power trip when on the Lamp | Emergency Lamp over flow Amperage | System Check, battery blow. | Problem Solved by Pembangunan. | CLOSED | HASRUL
17 | 2012-02-15 | Smoke detector | Alarm | False Alarm | Shortage. Pending vendor to check. | To contact Borhan (vendor). Suspect Battery run out. To change the battery. | CLOSED | MIRA
18 | 2012-03-05 | List Shortage Gas | Shortage | Shortage | To Order Spare | N2 for Oxidation Furnace, found spare tank | CLOSED | HASRUL
[6 empty] | SF6 for ICP, pending to contact vendor | CLOSED | ISA/MIRA
Pending quotation. To source to Indah.
19 | 2012-03-06 | Mask Aligner | Chrome Mask break when using smaller sample | When perss Vac Con | Use 4 inch for dummy | Problem Solved by technical team | CLOSED | HASRUL
20 | 2012-03-07 | Mask Aligner | No laser for align | No laser for alignment |  | Contacted vendor, pending Korean feedback. Laser Board courrier from Korea, pending to arrive and install by Nanorian. Wrong parts. Reorder 29 May. LED problem. Reinstalled independent power supply board on the 27th July. | CLOSED | HASRUL
21 | 2012-03-08 | Silent Compresor | fuse blown | fius blown in the middle of the process. | To comm with Magna value Vendor | To contact nazri (vendor). Suspect motor problem. Temporary used spare chiller from Magna Value. WW11 | CLOSED | HASRUL
22 | 2012-03-26 | RTA | No space for RTA |  | Waiting vendor coming to pair the new cleanbooth | To check the Gas, required N2 purge Gas, O2 99.9998% | CLOSED | JASNI
Pending final power supply.
All facilities ready, pending Hook up 25 July. Completed Hook up and commissioned on the 27th July.
23 | 2012-03-30 | Spin Coater #2 | Motor damaged | Motor can not runing | To comm with Nanorian Vendor | To contact Herman (vendor). Suspect motor problem. WW14, pending to check the motor from Ban Edar Sdn Bhd. Rejected. Sent to nanorian. | CLOSED | HASRUL
24 | 2012-04-02 |  |  | Pass down from Emi | Isa to compile the list | Pending Srubber Cert. No Passdown List. | CLOSED | ISA/EMI
25 | 2012-04-02 | Website |  |  |  | Pending list from all. Progressing. Thursday presentations 24 May. | CLOSED | ISA
26 | 2012-04-17 | Chemical storage | Not enough storage |  | Move to other room | Pending Foo to move out from Chemical storage room. | CLOSED | MIRA/Hasrul
|  |  | No Exhaust |  | TBD | Pending Sulaiman Quotation. | OPEN | JASNI
|  |  | No MSDS |  | NO MSDS, out from the Lab | To isolate and disposed from Lab by Oct | CLOSED | MIRA/Hasrul
27 | 2012-05-07 | Clear Gas tank | Space constraint | NO ID | To recycle the tank and source out to Indra Indah. | WW21 | CLOSED | JASNI
28 | 2012-05-08 | Chemical Database | No Control on access | No Tracibility | Transfer all Chemical to Chemical Room | Progressing on Data Base. | CLOSED | MIRA
29 | 2012-05-08 | Aircond Line | Water Leak | Danger to Power supply | Cover with Napkin | Forward complaint to Pembangunan. Remove the trunking, no more water drip. | CLOSED | Jasni
30 | 2012-05-08 | More Power Supply | Not enough socket | Machine swap power | Swapping Power socket | Pending pembangunan. 30th May to refurnish the power socket. 1 Ready. Pending for RTA, Approval from PM Mazelan. | CLOSED | Jasni
31 | 2012-05-29 | Chiller LPCVD tagging | No Tagging | Tagging rusted. | Request from Bendahari new Tag | Received on 28th May | CLOSED | Jasni
32 | 2012-05-25 | LT Furnace | Fuming haywire | Fuming haywire | Connecting the exhaust to Scrubber | Engage Contractor Sulaiman June | PROGRESS | Hasrul
33 | 2012-05-25 | Convection Oven | Fuming haywire | Fuming haywire | Connecting the exhaust to Scrubber | Engage Contractor Sulaiman June | PROGRESS | Hasrul
34 | 2012-05-25 | Ducting Oxide furnace | Fuming haywire | Fuming haywire | Connecting the exhaust to Scrubber | Engage Contractor Sulaiman June | PROGRESS | Hasrul
35 | 2012-05-25 | Chemical exhaust | Fuming haywire | Fuming haywire | Connecting the exhaust to Scrubber | Engage Contractor Sulaiman June | PROGRESS | Hasrul
36 | 2012-05-25 | Emergency door | No Exit | No Exit | Re-do the door | Engage Contractor Sulaiman June | PROGRESS | Hasrul
37 | 2012-05-18 | Equipment Labelling | No ID | No ID | Re-lable | Re-lable | CLOSED | MIRA
38 | 2012-05-18 | Equipment explanation | No explaination | No explaination | To stick explaination | Re-stick | CLOSED | MIRA
39 | 2012-05-18 | Lab coat | No tagging/ownership | No tagging/ownership | To buy extra, Will do the assignment system. | Order new set as a back up. 2013 to re open. New set ordered. | CLOSED | MIRA
40 | 2012-07-25 | 5's Housekeeping | Lack of order | No checking | To conduct Housekeeping | 1 week. Housekeeping activity was carried out on the 2nd Aug 2012. Tq all. | CLOSED | JASNI
41 | 2012-08-27 | Auto clave commissioning | Pending certifications | Pending FC Bios arrangement with JKKP Perlis | Between 17 th to 21st Sept. | Completed on the 20th Sept 2012 | CLOSED | JASNI
42 | 2012-08-27 | LPCVD Recovery | Process escaped pressure limit | Process for deposition poly silicon cannot run | Waiting vendor Johari coming | WW35 | CLOSED | MIRA
43 | 2012-08-27 | Evaporator Turbo Pump error | Turbo Pump fail | Turbo Pump fail | need to servis | Service vacuum line completed. 24th Sept 2012 | CLOSED | Hasrul
44 | 2012-08-27 | Lab Modifications paperwork/lab facilities improvement. | Safety Hazard | No proper Safety Spec | To re-engage contractor for Quotation | Abort due to budget not approved. | ABORT | Isa
44 | 2012-08-27 | Chemical Waste Pump | Pump kaput | Cannot evacuate Chemical Waste | To refer to Vendor that Supply the pump | WW35 | CLOSED | HASRUL
45 | 2012-09-04 | Fume Hood Air flow/hygiene tester | Air Flow Test | Air Flow Test | To conduct Air flow test and maintain record. | Tester to be provided by UKKP. Completed and passed 21st Sept 2012 | CLOSED | JASNI/ISHAK JANOO
46 | 2012-09-28 | Gravity/Natural Convection Oven | fuse blown | Current overflow | DOWN | Suspect water cause shortage. Repair by vendor on 23 Nov | CLOSED | HASRUL
47 | 2012-10-11 | Evaporator heater cable burnt | burnt smell and smoke | heater cable burnt | replace heater cable and reconnect power line | separate the cable and heater | CLOSED | HASRUL/JASNI
48 | 2012-10-11 | Budget for 2013 | All | All | All | WW51. Completed 17 Dec 2012 | CLOSED | JASNI
49 | 2012-10-11 | Gas tank change Bann Edar to Global Exergy | Change vendor | Change vendor | Change vendor | To discuss xfer period. 17 Dec 2012, only 2 tanks to keep. CF4 & SF6 | CLOSED | ISA/HASRUL
[5 empty] | Abort | Abort the process due to too expensive (near 50k). Continue with Bann Edar while refilling to be sun-con.
50 | 2012-11-26 | 5's Housekeeping | INEE backside lab mess | INEE backside lab mess | to organize housekeeping 3 Dec 2012 | Organizing housekeeping on the every Friday 3-5pm | CLOSED | All
51 | 2012-11-26 | Oxide Furnace | Ceramic crack | Tube over run | Temporary usage | Claiming under warranty from vendor. Ceramic still can be used. | CLOSED | HASRUL
52 | 2012-11-26 | Low power Scope | Component embark to Mask aligner project | Low Power Scope not functioning | To buy or replace | Isa to justify from Prof Uda. Due to low usage, will be obsoleted per the procedure. | CLOSED | ISA
53 | 2012-11-26 | Process Control | No Process Control | No Process Control | To hire intern to execute | Apply Intern to CiC, 2014 Project. | ABORT | ISA
54 | 2012-11-26 | Re-schedule Equipment usage | Too frequent | Too frequent | To reduce the frequent | 1 week | CLOSED | MIRA
55 | 2013-01-08 | Mask Aligner | Cannot focus | No light from source | Temporary using the manual Mask aligner | Pending Nanorian to order from Korea (Source Box). Completed 20th March. | CLOSED | Hasrul/Jasni
56 | 2013-01-16 | Evaporator | Titanium Dep cannot reach temperature | Cable connector burnt | Temporaraly using old cable | Need to use ceramic insulator at the connector point, to send to vendor for insulator seal. High insulator cable delivered on the 13th Feb 2013 | CLOSED | Jasni
57 | 2013-01-18 | Evaporator | Front shield not fit into the chamber | Dimension error | temporay remove the shield. | to be resent to the vendor for modifications. Parts received on 6th April | CLOSED | Jasni
58 | 2013-02-07 | LPCVD | LPCVD chiller problem | chiller not runing | to buy new motor chiller | pending the new motor chiller from shahrul. Motor replaced on the 13th Feb 2013. | CLOSED | Jasni | .
59 | 2013-01-18 | Silent Compresor | Compressor failed | No CDA | replaced PU tubing | Bought Tube from Bann Edar | CLOSED | Hasrul/Jasni
60 | 2013-02-27 | DNA Booth | RH% Failed | RH% requirement <50% | To install aircond | Abort the aircond installations due to lack of space. | CLOSED | Mira
61 | 2013-02-27 | DNA Booth | Need extra access | No room access | To open extra door curtain | 2 months (done 12 April) | CLOSED | Mira
62 | 2013-04-01 | LPCVD | 1-Thermocouple rear heater kaputt, 2-Temp ramp too slow, 3 to re-order 2nd boat | 1-Thermocouple rear heater kaputt reading only 592 instead of 600 and confirmed after swap with center thermocouple, 2-Temp ramp too slow due SSR kaputt or old single phase wiring, 3 to re-order 2nd boat | Hi-tech craft to conduct the audit | since Jan 2013, re-wiring to reconfirm mid May. Intevac attended the problem, found that maybe Xformer too old. Confirming the voltage to re-wiring. 10 July 2013. | CLOSED | Jasni
63 | 2013-05-02 | 3D Profiler | PC cannot ON & Hang. | Cable connector problem | waiting vendor from magna value | vendor coming at 27 may 2013. Change RAM Card on the Sept 18. | CLOSED | Mira
64 | 2013-05-24 | Mask Aligner | uv not good (blurred) | expose not uniform | pending uv from korea | Replace new Bulb. On hold due to power current no drop. Vendor attended the problem (Herman), extending the ON power to reactivate gas mercury. Problem solved. | CLOSED | Hasrul
65 | 2013-10-08 | LPCVD Recovery | Tube Evac error | Gas line pressure | Manual evac | To run cycle purge. Suspect the earlier setting was changed and causing the communication hang from the standby step. | CLOSED | Jasni
66 | 2013-10-15 | Evaporator | Minor hair leak | -3 torr minor leak | To change the O ring | Pending O ring re supply. Received O ring and changing the O ring. | CLOSED | Jasni
67 | 2013-10-09 | LPCVD | Silane gas flow stop | During Process step 2, Throttle valve stop to regulate causing process abort | Abort process, t/shoot the comm. | Suspect Thermocouple problem causing the comm abort. Replacing the part. Thermocouple returned on the 23 Dec 2013. | CLOSED | Jasni
|  | LPCVD | Silane gas flow stop | During Process step 2, Throttle valve stop to regulate causing process abort | Abort process, t/shoot the comm. | Increase error rate to 30%. Process ran as per normal. Completed 13 Jan 2014. | CLOSED | Jasni
68 | 2013-10-09 | RTO | Upgrading Process | Gauge Pressure showing 0.15 Torr instead of 760 Torr. | Doing off line gas monitoring , not calibration problem. MFC also doing off line for 20 SLM display | Both reading doing off line display. Change Pressure monitoring to Offline display. Completed on 7 Feb. Upgrade MFC to 20,000 sccm. RTO process upgrade completed 7 Feb 2014. | CLOSED | Jasni/Hasrul/Azman/Hafiza
69 | 2014-02-03 | ICP RIE | FR generator lamp kaput. | No error but for safety countermeasure. | Down equipment for switch replacement. | Switch stuck. Replaced on the 7 Feb 2014. | CLOSED | Hasrul.
|  |  | FR generator lamp kaput. | RF no plasma | Check continuity current power supply | Replaced 3A Fuse | CLOSED | Hasrul.
70 | 2014-02-03 | Muffler Furnace | Heating element kaput. | Error for heating limit. | Vendor Magna Value reset the setting. | Heating recovered. | CLOSED | Hasrul.
71 | 2014-02-03 | 3D Profiler | Measurement hang | During measurement process, software hang. | Abort, pending magna value vendor to attend the issue. | Only after 1st measurement is valid, the 2nd trial will be deviated. | PROGRESS | Mira.
72 | 2014-01-22 | Dieletric Analyzer | Measurement hang | During measurement process, software hang. | Abort, pending new GPID card ordering | Lead time vendor to replce GPID card 20th March. Not GPID problem. Motherboard kaputt and replaced. | CLOSED | Dr Foo./Jasni
73 | 2014-03-03 | PL | Extreme Peak calibration out | when doing high peak cal, reading uncertain | re-do calibration | Pending swap parts, software not compatible. Software updated | CLOSED | Kashif/Isa
74 | 2014-06-02 | Plasma preen | Wave fan not turning | during Micrwave activation, fan not turning | to open and re align the fan | pending open WIP, limit only for short time 2 minutes process | CLOSED | Mira/Jasni
75 | 2014-08-13 | UV Vis | lamp kaputt | lamp kaputt | replacement | pending replacement & Servicing, Equipment still can be used | CLOSED | Jasni/Mira
76 | 2014-08-13 | Waste pump | burnt and kaputt | 86l/m too low | to replace higher power pump | pending replacement. Pump serviced on th 25 Sept. resume back the normal process. | CLOSED | Jasni
77 | 2014-08-13 | Spin coater vacuum system | vacuum error | intermittent error | to re-layout the vacuum line and CDA line | Pending re-lay out. Re-layout completed on the 4th Nov by RA Azman Abu Hasan. | CLOSED | Jasni
78 | 2014-08-13 | Evaporator | vacuum error | intermittent error | gauge mesh cover | to fabricate the wire mesh, received the wire mesh on the Sept 23rd. Vacuum error reported due to less debris fall through the Vacuum Gauge thru put. | CLOSED | Jasni
79 | 2014-08-13 | Gas tank regulator | leak | gas leak | to replace regulator | to replace regulator. Pressure recovered by releasing the gas thru the back flow. Resume the process as per normal on the Oct 7th. | CLOSED | Jasni
80 | 2014-08-13 | I-V cable mess | magnetic field flux | Lay out mess | to re-layout the cable | to re-layout the cable. Pending Intern for deeper understanding. | CLOSED | Jasni/Azman
81 | 2014-08-13 | CVD FBR Hook up | Hook up list | Cost | vendor listing | Vendor listing Done, pending delivery in Feb 2015. Incoming power commando single phase completed on the 10th Jan 2015. | CLOSED | Jasni
82 | 2014-08-13
83

## PIC
INEE Lab Monitoring (Last updated 3 March 2015)
[12 empty] | PIC |  | vendor | name | contact no
|  |  | Status | Intro | SOP | Process Control | Maintenance | Safety | HIRARC | Safety Cert | Certification | 1st | 2nd
No | ID | Tool
1 | ICP RIE | Reactive ION Etching | UP | OK |  |  | OK |  | OK |  |  | AIZAT | MIRA | NANORIAN TECHNOLOGIES | Mohd Yazid Ahmad,Mohd Herman | 019-3997740, 0132417112 tel 03-87395090, yazid@nanoriantech.com
2 | RTA/RTO | Repeat Thermal Anneal/ Repeat Thermal Oxidation | UP | OK |  |  | OK |  |  |  |  | AIZAT | MIRA | NANORIAN TECHNOLOGIES | Mohd Yazid Ahmad,Mohd Herman | 019-3997740, 0132417112 tel 03-87395090, yazid@nanoriantech.com
3 | LPCVD | Low Pressure Chemical Vapour Dep | UP | OK | OK |  | OK | I/O Button,Emergency Button,EXH Fan Unit Suis,Heater Power Suis,MCB LPCVD Transformer, Heater CB, Main System Power. | - |  |  | AIZAT | MIRA | ULVAC | Shahrul, Hisham | 012-423-5157, 019-411-7219
4 | HT Eva | Thermal Evaporator | UP | OK | OK |  | OK | I/O Button,Reset Button,Cooling Water Supply | OK |  |  | MIRA | AIZAT | ULVAC | Shahrul, Hisham | 012-423-5157, 019-411-7219
5 | HP | Hot Plate | UP | OK | OK |  | N/A | I/O Button | OK |  |  | MIRA | AIZAT
6 | SC | Spin Coater 1/2/3 | UP | OK | OK |  | OK | I/O Button | N/A |  |  | MIRA | AIZAT | NANORIAN TECHNOLOGIES | Mohd Yazid Ahmad,Mohd Herman | 019-3997740, 0132417112 tel 03-87395090, yazid@nanoriantech.com
7 | SP | Spinner | UP | OK | OK |  | - | - | - |  |  |  | AIZAT | AdHOC
8 | USC | Ultra sonic Cleaner | UP | OK | OK |  | N/A | I/O Button | N/A |  |  | MIRA | AIZAT | AdHOC
9 | LT F | LT Furnace | UP | OK | OK |  | OK | I/O Button, Heater Button | OK |  |  | MIRA | AIZAT
10 | Vac Ov | Vacuum Oven | UP | OK | OK |  | N/A | I/O Button | OK |  |  | AIZAT | MIRA
11 | PPS | Plasma Preen System | UP | OK | OK |  | OK | I/O Button | - |  |  | MIRA | AIZAT | ULVAC | Shahrul, Hisham | 012-423-5157, 019-411-7219
12 | MA | Mask Aligner | UP | OK | OK |  | OK | I/O Button | OK |  |  | AIZAT | MIRA | NANORIAN TECHNOLOGIES | Mohd Yazid Ahmad,Mohd Herman | 019-3997740, 0132417112 tel 03-87395090, yazid@nanoriantech.com
13 | Ox F | Wet/Dry Oxidation Furnace | UP | OK | OK |  | OK | I/O Button, DOL for EXH Fan,Power ON Reset,Excess Temp Controller, Emergency Botton | OK |  |  | AIZAT | MIRA | HITECH Craft (M) Sdn Bhd | Johari | 016-4454272
14 | MUF F | Muffle Furnace | UP | OK |  |  | N/A | I/O Button, ON/OFF Unit Suis | OK |  |  | MIRA | AIZAT
15 | PS | DC Power Supply | UP | OK |  |  | N/A | MCB Point | OK |  |  | MIRA | AIZAT | AdHOC
16 | A Com | Air Compressor | UP | - | - |  | OK | I/O Button | N/A |  |  | AIZAT | MIRA | AdHOC
17 | Si Com | Silent/Oil Compressor | UP | - | - |  | OK | I/O Button | N/A |  |  | AIZAT | MIRA | Magna Value SDN BHD | Rosli Othman , Operation Manager | 013-468-5123, 04-425-7986, rosli.mvsb@gmail.com; SHUKOR (0194881869); NAZRI (0194881867)
18 | Con Ov | Gravity/Natural Convection Oven | UP |  |  |  | N/A | I/O Button, Hi Limit Selection | - |  |  | AIZAT | MIRA | Tech lab manufacturing sdn bhd |  | 03-89483300 email;;techlab@tm.net.my
19 | Cent | 500g Centrifuge | UP | OK | - |  | N/A | I/O Button | N/A |  |  | MIRA | AIZAT
20 | VO Mix | Vortex Mixer | UP | OK | - |  | N/A | I/O Button | N/A |  |  | MIRA | AIZAT
21 | Bal | Balancing | UP | OK | - |  | N/A | I/O Button | N/A |  |  | MIRA | AIZAT
22 | USC | ultra sonic cleaner DC200H | UP | OK | - |  | N/A | I/O Button | N/A |  |  | MIRA | AIZAT
23 | A Clav | AutoClave | UP | OK |  |  | OK | Safety valve | OK |  |  | MIRA | AIZAT | FC Bios sdn bhd | George Lim | 019-214-0853 John Benedict (john_benedict@fcbios.com.my), 019-2606520, 03-56351559 ( fara) ex:204 JKKP Perlis Tech Azizul 019-301-4556
24 | Cent | Centrifuge 5430/5430 R | UP | OK | - |  | N/A | I/O Button | N/A |  |  | AIZAT | MIRA | Medigene sdn bhd | Liang Yee Kai | 012-6469285, 1-300 88 8668, liang.yeekai@mdgsb.com
25 | PL | Photo Luminecence | [9 empty] | ISA | LIU | ALV Technologies Sdn. Bhd. | Lia Ahmad Ramli | 012-3213 707 malaysia@analytical-online.com 03-470 5386 Fax: 03-4270 4386
26 | SEM | JEOL SEM-EDS 6010LV |  |  | OK | N/A | Oil pump, filament change | I/O button | YES | Radiation Dept |  | ISA | Dr Foo | JEOL (M) SDN BHD | Peter & Ong | Peter 012-498-9592 (email peterchuah@jeolmal.com), Ong 012-427-6800 email (hcong@jeolmal.com), No 58, 1st Floor, Persiaran Mahsuri 1/2, Sunway Tunas, Bayan Lepas 11900, Penang, Malaysia Tel 04-645-1608.
27 | Constance Oven | Constance Oven | UP | TBA | TBA | N/A | TBA | TBA | High Temp |  |  | Dr Foo | Jasni | IRAMA CANGGIH SDN BHD |  | NO.7, LINTASAN PERAJURIT 17A, TAMAN PERINDUSTRIAN IPOH, 31400 IPOH, PERAK DARUL RIDZUAN. contact - 05-5487822 email - sales@iramac.net
28 | Thermo shaker | Thermo shaker | UP (20 ml & 24 ml) | TBA | TBA | N/A | TBA | TBA | Vibration, Voltage Current |  |  | Dr Foo | Jasni | Fisher Scientific (M) Sdn. Bhd. |  | 30, Lorong Helang 2, Desa Permai Indah, 11700 Pulau Pinang, Malaysia. contact - 6(04)-660 4100 email - Enquiry.my@thermofisher.com
No | ID | Analysis | UP
1 | Four Point Probe | Four Point Probe | UP | OK | OK |  | OK | I/O Button | N/A |  |  | ISA | AIZAT | eila sdn bhd | Alan Goii | Alan 012-4770201. Office 04-537-9340 add Lot 19, Lorong Janggus 5, Makok Garden 14000 Bt Mertajam, Penang Malaysia
2 | Osc | PC Oscilloscope | To be obsoleted. | OK | OK |  | N/A just for calibration | I/O Button | N/A |  |  | ISA | AIZAT
3 | Osc | Oscilloscope | To be obsoleted. | OK | OK |  | N/A just for calibration | I/O Button | N/A |  |  | ISA | AIZAT
4 | Hi-Pow | High Power Microscope 100X | UP | OK | OK |  | OK | I/O Button | N/A |  |  | ISA | MIRA | Crest Systems (M) Sdn. Bhd. | Victor Tang Snr Technical Support Engineer | H/P: +6012-425 7056 Office: +604-645 2268 (ext 103)
5 | Lo-Pow | Low Power Microscope 4.5X | To be obsoleted. | OK | OK |  | Not Yet | I/O Button | N/A |  |  | ISA |  | Check status
6 | 3D Profiler | 3D Surface Profilometer | UP | OK | OK |  | N/A just for calibration | I/O Button | OK |  |  | ISA | AIZAT | Magna Value SDN BHD | Rosli Othman , Operation Manager | 013-468-5123, 04-425-7986, rosli.mvsb@gmail.com; SHUKOR (0194881869); NAZRI (0194881867)
7 | Die-analyzer | Dielectric Analyzer | UP | OK | OK |  | OK | I/O Button | N/A |  |  | ISA | AIZAT
8 | Sig-Gen | Signal Generator | UP | - |  |  | N/A just for calibration | I/O Button | N/A |  |  | ISA | AIZAT
9 | Elli | Ellipsometer | Software NA. To be obsoleted. | TBA | TBA | TBA | Laser Calibration | Laser Exposure | Laser Exposure |  |  | ISA | MIRA | ULVAC | Shahrul, Hisham | 012-423-5157, 019-411-7219
10 | UV-IS | UV-IS Perkin Elmer UV/VIS Lambda 35 | UP | TBA | TBA | TBA | [5 empty] | ISA |  | D.I. SCIENTIFIC SDN BHD, no17 Jln SS7/26 Kelana Jaya 47301 Petaling Jaya Selangor | Munirah | 016-211-2405, Anchee 019-621-3828 (a lady)
11 | PicoAmp | PicoAmpere | UP, Pending Interface software | TBA | TBA | TBA | [5 empty] | ISA |  | NANORIAN TECHNOLOGIES | Mohd Yazid Ahmad,Mohd Herman | 019-3997740, 0132417112 tel 03-87395090, yazid@nanoriantech.com
12 | Filmetric | Filmetric | UP | TBA | TBA | TBA | TBA | TBA | Laser Exposure |  |  | ISA |  | Kamuning | Jamil | hp 016-411-5766, off 04-507-2696,04-507-5969; hazlina@kemuningsaintifik.com (Hazlina: 60-12-5793767)
13 | GAMRY | Potentiostat | UP | TBA | TBA | N/A | TBA | TBA | Voltage current |  |  | Dr Nazree | ISA
14 | LCR | LCR Meter | UP | TBA | TBA | N/A | TBA | TBA | Voltage current |  |  | Dr Nazree | ISA
No | ID | Facility
1 |  | White Clean room booth | UP | OK |  |  | OK | I/O Button for lamp and fan | N/A |  |  | MIRA | AIZAT
2 |  | Yellow Clean room booth | UP | - |  |  | OK | I/O Button for lamp and fan | N/A |  |  | MIRA | AIZAT
3 |  | DI Water | UP | - |  |  | N/A | I/O Button | N/A |  |  | AIZAT | MIRA | Adscitech (M) SND.BHD | Che Umar | 019-4578877, 04-4255758 omaq228@yahoo.com
4 |  | Chemical Bath | UP | - |  |  | OK | N/A | OK |  |  | AIZAT | MIRA
5 |  | Fume Hood in yellow booth | UP |  |  |  | OK | I/O Button | N/A |  |  | AIZAT | MIRA
6 |  | Electronic Desiccators | UP | OK |  |  | -
7 |  | Micro-pipette | UP | OK |  |  | N/A | N/A | N/A |  |  | MIRA | AIZAT
8 |  | Scrubber System | UP |  |  |  | N/A | N/A | OK |  |  | AIZAT | MIRA
9 |  | chemical storage | UP |  |  |  | N/A | N/A | OK |  |  | MIRA | AIZAT
10 |  | fume hood in nano material area | UP | OK |  | Flow Test | OK | I/O Button | N/A |  |  | AIZAT | MIRA
11 |  | Low Temprature Freezer | UP |  |  |  | N/A | I/O Button | OK |  |  | MIRA | AIZAT | Adscitech (M) SND.BHD | Che Umar | 019-4578877, 04-4255758 omaq228@yahoo.com
12 |  | Clarke | UP | OK |  |  | N/A | I/O Button | OK |  |  | AIZAT | MIRA
13 |  | Freezer | UP | OK |  |  | N/A | I/O Button | OK |  |  | AIZAT | MIRA
14 |  | DNA Laminar Flow Cabinet | UP | OK |  |  | OK | I/O, filters | progress |  |  | AIZAT | MIRA | TEAM MEDICAL & SCIENTIFIC SDN BHD Address No. 41, Jalan Anggerik Vanilla T31/T, Kota Kemuning, 40460 Shah Alam, Selangor, Malaysia |  | Phone No. +603 5638 0348 ¦ +6019 - 342 1752
15 |  | Gloves Box | UP | OK | TBA | N/A | N/A | View Mirror & Gloves no leak | No Leak from Inside to outside | N/A | N/A | AIZAT | MIRA | N/A | N/A | N/A
16 |  | Water Pump Filter (building) | UP |  |  |  | OK, Weekly |  |  |  |  | AIZAT | MIRA
17
18
No | ID | Supporting
1 |  | Chemical Handling |  |  |  |  | N/A |  |  |  |  | MIRA | AIZAT | Adscitech (M) SND.BHD | Che Umar | 019-4578877, 04-4255758 omaq228@yahoo.com
2 |  | Chemical Register |  |  |  |  | N/A |  |  |  |  | MIRA | AIZAT
3 |  | Gas Room |  |  |  |  | N/A |  | NO ENTRY |  |  | AIZAT | MIRA
4 |  | Chemical Cabinet |  |  |  |  | N/A | Scrubber System | Dangerous area |  |  | AIZAT | MIRA
Note on F6: mira: Hook up completed on the 20th July 2012
Note on F7: mira: 18 May 2012 since April 2011.
Note on F44: mira: UP ( 10 aug 2012 )
Note on F50: mira: Comissioned on the 25th Feb 2013

## Lab Working Sheet & Results
Lab Technician Working Sheet 2013 and results
Seq | Date | PIC | Sample owner | Project | Equipment ID | Task & Process | Results | Remark
1 | 2013-01-30 | Hasrul | Adam | Nano Wire | ICP | Poli Silicon etching = recipe 18 (T=1min)
Resist etching = recipe 22 ( 3min,5min)
si etching = recipe 5 (1min)
2 | 2013-02-14 | Hasrul | Nuzaihan/FYP Vera/sharma/shahrulAizam | Si nano wire | LPCVD | Si=1200secs, Gas Si=80sccm, Temp 600 ALL (23 samples) (Wafer sequence ) (Rear)/(Nuzaihan 3-6, FYP 7-16, Vera 17-20, Azam 21-22)
3 | 2013-02-14 | Hasrul | Nuzaihan |  | Oxide Furnace (Wet) | N2=10 l/m, O2=15 l/m, 1 hr duration | Thickness=260 nm, U%=?
4 | 25-Mac-13 | Mira | Nuzaihan |  | Oxide Furnace (Wet) | N2=10 l/m, O2=15 l/m, 1 hr duration
5 | 26-Mac-13 | Mira | En Ramzan, Nan |  | LPCVD | Si=1200 secs, Gas Si=80 sccm, Temp 600 All (8 sample)/ En Ramzan 3 qty (22-24) Nan 3 qty (18-20) , Mira 1 qty (21)
6
7
8
9
10

## Training
INEE Training Record Sheet
Date | PIC | Student Name | Sample owner | Project | Eqipment ID | Assignment | Period | Time taken (Hrs)
2012-01-20 | Mira/Hasrul | Syarifah ,Adam,Kashif | NA | Training | 3D Profiler | Certifications | 10am~1pm | 3

## Gas list
INEE Lab Gas List (Updated 22 DEC 2014)
Area | Tool ID | Tank # | Gas Name | Purity | Tank no | TBD | Symbol | Status | Date Refill | Level (Pressure) | Vendor | Contact | To buy? | Gas Tank no | Date in | Date Out
Material Lab | LT Furnace | 1 | Argon Compressed (Big) | 99.8 | NO:137905 |  | Ar | Active-dr nazri | 24/06/2014 | Full |  |  | No
|  | 2 | Oxygen | 99.8 | NO:NKH003717 |  | O2 | Active-NurHamidah | 16/07/2012 | Full | Ever point | 019-4578877, 04-4255758 omaq228@yahoo.com | No
[12 empty] | Suhaizat | No
Material Lab | Gas Sensing Setup | 3 | Purified Nitrogen | 99.999 | NO:A 65964 |  | N2 | Active-Kashif | unknown | Full | Bann Edar | 04-9765570 | No
|  | 4 | Purified Air (small) | 99.999 | NO:105W14 |  | AIR | Active-Kashif | unknown | Full |  |  | No
|  | 5 | Purified Air (big) | 99.999 | NO:117351 |  | AIR | Active-Kashif | 2009-03-12 | Full |  |  | No
|  | 6 | 1% Diluted Hydrogen | 1% H2 | NO:405586 |  | H2 | Active-Kashif | 21/10/2011 | Full |  |  | No
Gas Room (Back) | Oxidation Furnace | 22 | Oxygen | 99.8% | NO:4003 |  | O2 | Active | 2014-07-06 | Full | Che Umar | 019-4578877, 04-4255758 , 04-425-1204,omaq228@yahoo.com | No
|  | 35 | Nitrogen (Purge) | 99.8% | NO:06650 |  | N2 | Active | 2014-07-06 | Full |  |  | No
| Spare | 23 | Oxygen | 99.8% | NO:2010 | 99.8% | O2 | Finish | 25/4/2014 | to be refill | Bann Edar/ ever point | 04-9765570 | Yes
|  | 7 | Nitrogen (Purge) | 99.8% | NO:001 | 99.8% | N2 | Finish | 25/4/2014 | to be refill |  |  | Yes
|  | 34 | Nitrogen | 99.8% | NO:89310 | 99.8% | N2 | Finish | unknown | to be refill |  |  | Yes
|  | 27 | Nitrogen (small) | 99.8% | NO:1006 | 100% | N2 | Standby | 19/10/2011 | Full |  |  | No
|  | Missing | Argon (Small) | 99.8% | NO:505630 |  | AG | Standby | 28/11/2011 | Full |  |  | No
|  | 26 | Sulfur hexafluoride (Small) | 100% | NO:17770 |  | SF6 | Standby but no need to use | 25/12/2012 | Full |  |  | No
|  | 31 | Purified Argon | 99.99% | NO:12756 |  | AG | Standby | 2013-01-03 | Full |  |  | No
|  | 33 | Carbon tetrafluoride (big) | 100% | NO:22033 |  | CF4 | Standby | 2013-01-03 | Full |  |  | No
|  | 30 | Argon (Small) | 99.8% | NO:505630 | 99.8% | N2 | Standby | 28/4/2011 | Full |  |  | No
|  | 24 | Oxygen | 99.8% | NO:137404 |  | O2 | Standby | 25/4/2014 | Full |  |  | No
|  | 29 | 1% H2 BALN2 (small) | 1% H2 | NO:475853 |  | H2 | Standby | 2012-10-10 | Full |  |  | No
|  | 28 | Helium | 100% | NO:118609 |  | He | Standby | 2013-02-04 | Full |  |  | No
|  | 32 | Argon Compressed (Big) | 99.8 | NO:72903 |  | Ar | Finish | unknown | to be refill
|  | 25 | Nitrogen Pure | 100% | NO:SIG 1441 | 100% | N2 | No Status ( red tank) | unknown | Full |  |  | No
|  |  | LPG ( Gas Petronas) | NA |  | NA | NA | Active Kasyif | unknown | Full |  |  | No
Gas Room (Back) | ICP/RIE | 14 | purified air | O2 21% | NO:93886 |  | Ar | Active | 2011-12-09 | Full | Bann Edar | 04-9765570 | No
|  | 15 | Nitrogen Purge | 100% | NO:458-311-00343632 |  | N2 | Active | 28/3/2014 | Full |  |  | No
|  | 16 | Helium (Small) | 100% | NO:8066 |  | He | Active | unknown | Full |  |  | No
|  | 17 | Purified Argon (Small) | 100% | NO: 3259 |  | Ar | Active | unknown | Full |  |  | No
|  | 18 | Trifluoromethane (Small) Silicon Oxide+ Silicon Nitride | 100% | NO:930227 |  | CHF3 | Active | unknown | Full |  |  | No
|  | 19 | Oxygen | 100% | NO: A0297 |  | O2 | Active | 28/5/2014 | Full |  |  | No
|  | 20 | Carbon tetrafluoride (Small) Silicon Oxide | 100% | NO:831707 |  | CF4 | Active | unknown | Full |  |  | No
|  | 21 | Sulfur hexafluoride (Big) Polysilicon | 100% | NO:9958725 |  | SF6 | Active | New 99.993% 12 July 2012 | Full |  |  | No
No
Gas Room (Inside) | LPCVD | 13 | Silane | 99% | NO:MOX 150708 |  | SiH4 | Active | 17/5/2010 | Full | Bann Edar | 04-9765570 | No
|  | 11 | Nitrogen Vent | 100% | NO:4250113 |  | N2 Vent | Active | 30/1/2013 | Full |  |  | No
|  | 12 | Nitrogen Purge | 100% | NO:5140513 |  | N2 Purge | Active | 20/5/2013 | Full |  |  | No
Nano Biochip 2 | Plasma Preen System | 8 | Purified Oxygen | 99.8% | NO:66571 |  | O2 | Active | 26/3/2013 | Full | Bann Edar | 04-9765570 | No
|  | 9 | Argon (Small) | 100% | NO:D7020973 |  | Ar | Active-Never Used | unknown | Full |  |  | No
Nano Biochip 2 | RTA | 10 | Nitrogen (Purge) | 99.8% | NO:901292 |  | N2 | Active | 2014-01-05 | Full | Che Umar | 019-4578877, 04-4255758 omaq228@yahoo.com | No
|  | 36 | Nitrogen (Purge) | 99.8% | NO: 76706 |  | N2 | Active | 2014-12-02 | Full | Bann Edar | 04-9765570 | No
.
TOTAL = 36 TANK

## Equipment Schedule Dec 2012
INEE LAB MAJOR EQUIPMENT OPERATION SCHEDULE
UPDATED 17 DEC 2014
STARTING 9.00 AM - 4.30 PM
MONDAY | LPCVD / MASK ALIGNER
TUESDAY | THERMAL EVAPORATOR / OXIDATION FURNACE
WEDNESDAY | THERMAL EVAPORATOR / MASK ALIGNER / ICP-RIE
THURSDAY | OXIDATION FURNACE / MASK ALIGNER
FRIDAY | ICP - RIE ( STARTING 9.00AM-12.00PM)
Technician and PLV in-charge contact #:
PN MIRA (Technician INEE) | : 017-466-1969
EN JASNI (PLV INEE) | : 019-411-8081
FOR SEM/EDS & PL MINIMUM 1 WEEK BEFORE BOOKING
PLEASE CONTACT EN MOHAMMAD ISA 013-4888973 FOR BOOKING
Please pre-book your schedule with the Technician in-charge. Any urgent request is subjected to approval by INEE's Director.

## ICP Recipe
ICP RIE Recipe Setting
MATERIAL | Recipe # | CF4 (sccm) | CHF3 (sccm) | SF6 (sccm) | O2 (sccm) | Ar (sccm) | Bias (V) | ICP Power (W) | APC/Control Pressure | Time recommended? (sec) | Etch Rate? (Å/sec) | Thickness (Å)
Silicon | 5 | 30 | 0 | 0 | 28 | 0 | 300 | 500 | 5 | 30~180
Silicon Nitride | 13 | 0 | 12 | 0 | 0 | 30 | 250 | 800 | 2.5
Silicon Oxynitride | 20 | 0 | 12 | 0 | 0 | 30 | 250 | 800 | 1 |  |  |  |  | 709.2 | 23.64
Silicon Oxynitride | 14 | 0 | 12 | 0 | 0 | 30 | 250 | 800 | 8 |  |  |  |  | 711.6 | 23.72
Silicon Oxide (Thin) | 15 | 50 | 0 | 0 | 0 | 30 | 250 | 800 | 5 |  |  |  |  | 725.6 | 24.19
Silicon Oxide (Very Thin) | 15 (a) | 50 | 0 | 0 | 0 | 30 | 250 | 100 | 5 | 30 secs | 23.85 |  | Ave | 715.4667 | 23.85
Silicon Oxide (Very Thin) | 15 (b) | 50 | 0 | 0 | 0 | 30 | 250 | 100 | 5 | (10 + 10) secs | 23.4667
Silicon Oxide (Thick) | 6 | 50 | 0 | 0 | 0 | 30 | 300 | 800 | 5 |  |  |  |  | 474.4 | 23.72
Si Resist P. Step1 | 22 | 10 | 15 | 0 | 0 | 10 | 250 | 800 | 5 | Release back to production Dec 2012 |  |  |  | 463.1 | 23.16
P. Step2 |  | 0 | 0 | 0 | 20 | 0 | 30 | 200 | 1 | Release back to production Dec 2012 |  |  |  | 470.3 | 23.52
Remove Resist Only |  | 0 | 0 | 0 | 20 | 0 | 30 | 200 | 1 |  |  |  | Ave | 469.2667 | 23.4667
Poli Silicon | 18 | 0 | 0 | 50 | 0 | 30 | 250 | 650 | 4 | 300 (BKM)
Amorphous silicon | 19 | 50 | 0 | 0 | 0 | 30 | 300 | 500-700 | 5
Minimum Etching Effect |  | 50 | 0 | 0 | 0 | 30 | 50 | - | 5
Revision 1, 20 June 2012
Revision 2, 13 Oct 2014
Note on L22: ismail - [2010]:

## Lab user Reference number
LAST UPDATE 030315
INEE-sequence-DDMMYY
Item | User Name | FYP/MSc/PhD | Contact no | Origin/School | Equipment | Supervisor | reference Number
1 | Alaba Arauyinbo | [5 empty] | INEE 000001-130912
2 | Safaa Idan Mohammed | [5 empty] | INEE 000002-221012
3 | Husnee R. Abd | [5 empty] | INEE 000003-221012
4 | DR Sabri M. Hussein | [5 empty] | INEE 000004-261012
5 | DR Rejap Mohammed Hussen | [5 empty] | INEE 000005-261012
6 | DR Hamid S. Al-Jumaili | [5 empty] | INEE 000006-261012
7 | Mustafa Adil Abdul Rahim | [5 empty] | INEE 000007-261012
8 | DR Jihad Abed Taees | [5 empty] | INEE 000008-261012
9 | DR Ahmed Sarhan Mohammed | [5 empty] | INEE 000009-261012
10 | DR Ismail K. Ibrahim ( AL-Khateeb) | [5 empty] | INEE 0000010-261012
11 | DR Jamal Hameed Waheb | [5 empty] | INEE 0000011-261012
12 | KHALID AZADI BIN MAT AKHIR | [5 empty] | INEE 0000012-261112
13 | FASAL HAMAD | [5 empty] | INEE 0000013-291112
14 | HAMEED HUSSEIN ALI | [5 empty] | INEE 0000014-291112
15 | SADDA ABED ABDUALLAH | [5 empty] | INEE 0000015-291112
16 | AHMED TELANT HAMMEEDI | [5 empty] | INEE 0000016-291112
17 | NABEEL ARIF TAWFEEQ | [5 empty] | INEE 0000017-291112
18 | KALEID WALEED ABID | [5 empty] | INEE 0000018-291112
19 | IBRAHIEM JASSIM ABDULLEH | [5 empty] | INEE 0000019-291112
20 | DR ANMAR DHERAJ KOSAJ | [5 empty] | INEE 0000020-291112
21 | HARITH KAMIL BUNIYA | [5 empty] | INEE 0000021-291112
22 | RASHID MOHAMMAD RASHID | [5 empty] | INEE 0000022-291112
23 | MUHAMMAD SYAFIQ BIN SAMSUDIN | [5 empty] | INEE 0000023-131212
24 | MOHD ROSMADI BIN YAACOB | [5 empty] | INEE 0000024-131212
25 | ZULKHAIRI BIN ZAKARIA | [5 empty] | INEE 0000025-131212
26 | NURUL FARHANAH ABDUL HALIM | [5 empty] | INEE 0000026-131212
27 | KHOR KANG NAN | [5 empty] | INEE 0000027-131212
28 | MUHAMMAD HAISULHAFZAN BIN AHMAD | [5 empty] | INEE 0000028-131212
29 | NURUL MALIHAH BT MARZUAN | [5 empty] | INEE 0000029-131212
30 | FAREHANIM SYUKHAIRIYAH BT MOHD AKHIR | [5 empty] | INEE 0000030-310113
31 | NOR HUDA BT MOHAMAD | [5 empty] | INEE 0000031-310113
32 | MOHAMAD AZMAN BIN ABU HASSAN | [5 empty] | INEE 0000032-310113
33 | HASSAN AL FARUSH BIN BAKAR | [5 empty] | INEE 0000033-310113
34 | SITI NURHAFIEZAH BINTI ABDUL MUTALIB | [5 empty] | INEE 0000034-310113
35 | NURFARHANA AZIRA BINTI AHMAD IQBAL | [5 empty] | INEE 0000035-310113
36 | MUHAMMAD KAMIL BINWAN AHMAD | [5 empty] | INEE 0000036-070313
37 | CHEANG WEI HOW | [5 empty] | INEE 0000037-110313
38 | NURAINA BT MAINOR | [5 empty] | INEE 0000038-250313
39 | NEOH TZE TAI | [5 empty] | INEE 0000039-250313
40 | LIM WEE YANG | [5 empty] | INEE 0000040-250313
41 | CHAN BUAN FEI | [5 empty] | INEE 0000041-250313
42 | WONG WEE KEN | [5 empty] | INEE 0000042-250313
43 | MOHAMAD FARHAN ASRAF BIN MOHD ASLI | [5 empty] | INEE 0000043-170413
44 | LIM JOON HOONG | [5 empty] | INEE 0000044-170413
45 | ALI BASHEER AZEEZ | [5 empty] | INEE 0000045-170413
46 | ABDUL KADER ALAKRACH | [5 empty] | INEE 0000046-180413
47 | AYU WAZIRA AZHARI | [5 empty] | INEE 0000047-070513
48 | NURASMIN BINTI MOHAMAD RASLI | [5 empty] | INEE 0000048-230513
49 | NURAZLINA NAYAN | [5 empty] | INEE 0000049-230513
50 | NORASIKIN BINTI MUSTAFFA | [5 empty] | INEE 0000050-230513
51 | AHMAD HAMIDI BIN MOHD ZIN | [5 empty] | INEE 0000051-230513
52 | MUSARROFAH BT MOHD NOH | [5 empty] | INEE 0000052-230513
53 | NOOR SYAZWANIE BINTI MOHAMAD | [5 empty] | INEE 0000053-230513
54 | SURIATI BT SUHAIMI | [5 empty] | INEE 0000054-230513
55 | KHAIRIBASITAH BT KHAIRUDIN | [5 empty] | INEE 0000055-230513
56 | NAZATUL NADIA BT ABDUL HALIM | [5 empty] | INEE 0000056-230513
57 | NURUL FARHANAH ABDUL HALIM | MSc |  | School Bahan |  | Prof Mohd Nor/Dr Nazree | INEE 0000057-291013
58 | ABDUL RAHMAN BIN MOHD HADZIL | MSc |  | School Bahan |  | Dr Nazree | INEE 0000058-291013
59 | NORASFALIZA ABDULLAH | MSc |  | Sistem E |  | Dr Khor Shin Fhan | INEE 0000059-291013
60 | DEVI SHANTINI A/P CHANDIZAKARAN | PhD |  | School Bahan |  | Dr Irwana/Dr Nazree | INEE 0000060-291013
61 | ADELYN PUAH YING PING | FYP |  | Micro E |  | Prof Uda Hashim | INEE 0000061-131213
62 | STEPHEN SU CHIE SHANG | FYP |  | Micro E |  | Dr Mohd Khairuddin | INEE 0000062-181213
63 | YEE CHIEN CHIEM | FYP |  | Micro E |  | Dr Mohd Khairuddin | INEE 0000063-181213
64 | ANG YEW MING | FYP |  | Micro E |  | Dr Mohd Khairuddin | INEE 0000064-181213
65 | SAEED SALEM SAEED BASHASHWAN | FYP |  | Micro E |  | Dr Ruslinda A. Rahim | INEE 0000065-181213
66 | SYAFINAR BINTI RAMLI | MSc |  | System E | Only UV-IS | Gomesh Nair A/L Shasidharan | INEE 0000066-231213
67 | MUHAMMAD HAFIZ ABDUL RAHIM | MSc |  | Env ENG | Only Potentiostat | Dr Umi Fazara/Dr Nazree | INEE 0000067-301213
68 | NURSYAFIZA BT ISHAK | FYP |  | Micro E | SEM | Dr Ruslinda A. Rahim | INEE 0000068-080114
69 | TAN TIEK AUN | MSc |  | School Bahan | Only Potentiostat | Dr Sara Yasina/Dr Nazree | INEE 0000069-090114
70 | WONG HONG SEONG | FYP |  | MicroE | Only Evaporator | Mazlee/Nuzaihan | INEE 0000070-160114
71 | HADI MAHMODI SHEIKH SARMAST | PhD |  | USM School of Physics/IRAN | RTA | Prof MD. Roslan/Prof Uda | INEE 0000071-040214
72 | NITHIYAH A/P TAMILCHELVAN | FYP |  | Micro E | Spinner/Hot plate/Mask Aligner | Dr Shazlina/Nuzaihan | INEE 0000072-100214
73 | FATIN NABILAH MOHD FAUDZI | MSc |  | INEE | ALL | Dr Ruslinda A. Rahim | INEE 0000073-100214
74 | MAKRAM ABDUL MUTTALEB FAKHRI AL-JAFAR | PhD |  | INEE | ALL | PM Yarub Al-Douri | INEE 0000074-100214
75 | KWOK WAI KIT | FYP |  | MATERIAL/METAL | 4 POINT PROBE | NURULHUDA/DR NAZREE | INEE 0000075-060314
76 | NUR FATEHAH MOHD FAUZI | FYP |  | MATERIAL/METAL | TUBE FURNACE | DR ZURAIDAWANI/DR NAZREE | INEE 0000076-060314
77 | ANUL MADZIAH ABD WAHAB | FYP |  | MATERIAL/METAL | TUBE FURNACE | DR ZURAIDAWANI/DR NAZREE | INEE 0000077-060314
78 | KOH ZIE YANG | FYP |  | MATERIAL POLYMER | Only Potentiostat | JALILLAH JALIL | INEE 0000078-170314
79 | NISRIN BINTI ADLI | FYP |  | MATERIAL/METAL | Only 4 point probe | NURUL RAZLIANA/DR NAZREE | INEE 0000079-190314
80 | AHMAD FITRI BIN ANUAR MAHAYUDIN | MSc |  | MICRO E | LPCVD/3D PROFILO/DA | PM YUFRIDIN/DR KHAIRUDDIN | INEE 0000080-190314
81 | FAZMIR BIN HAMZAH | MSc |  | MICRO E | LPCVD/EVAPORATOR/YELLOW BOOTH/4 POINT PROBE | PM YUFRIDIN/DR KHAIRUDDIN | INEE 0000081-190314
82 | SINAR ARZURIA BINTI ADNAN | PhD |  | MATERIAL | 4 POINT PROBE | DR FIRUZ ZAINUDDIN/DR NAZREE | INEE 0000082-240314
83 | MOHAMMAD FIRDAUS (SAMPLES ONLY) | FYP |  | MICRO E | EVAPORATOR ONLY | MOHD ROSYDI/ DR DIN | INEE 0000083-080414
84 | ROSHIDA BINTI MUSTAFFA | PhD |  | MATERIAL | U-VIS | DR IRWANA/DR NAZREE | INEE 0000084-090414
85 | LEE KAN TAN | FYP |  | PHOTONIC ENGINEERING | PL | DR ALEDDIN A. SAIF | INEE 0000085-150414
86 | OOI ZEEN VEE | FYP |  | PHOTONIC ENGINEERING | PL | DR ALEDDIN A. SAIF | INEE 0000086-150414
87 | AHMAD MUHAIYAT JAMALUDIN | FYP |  | MATERIAL/METAL | 4 POINT PROBE | NURULHUDA/DR NAZREE | INEE 0000087-160414
88 | MANI RAJ NARTHAN | FYP |  | MATERIAL/METAL | OVEN | DR DEWI SURIYANI CHE HALIN | INEE 0000088-170414
89 | CONLATHAN A/L IBAU | FYP |  | MICRO E | ALL | PROF UDA | INEE 0000089-210414
90 | AHMAD LABIB BIN CHE AZIZ | FYP |  | MICRO E | ONLY EVAPORATOR | HARIYANTI/DR FOO | INEE 0000090-210414
91 | MUHSON NAJIY BIN ROSHIDEE | FYP |  | MICRO E | ONLY EVAPORATOR | HARIYANTI/DR FOO | INEE 0000091-210414
92 | SUSAN KAVITA MARTHAM | FYP |  | MICRO E | ONLY 3D PROFILER | DR MUKHZEER | INEE 0000092-250414
93 | NORSAIDAH BINTI MD ZUKI | FYP |  | BIOMEDICAL ELECTRONICS | ONLY 3D PROFILER | DR NASHRUL FAZLI MOHD NASIR | INEE 0000093-250414
94 | WAN MOHD HELMI HUSSEIN WAN OMAR | FYP |  | MATERIAL | 2 & 4 POINT PROBE | NURUL HUDA/DR NAZREE | INEE 0000094-250414
95 | NG SHIH MENG | FYP |  | MATERIAL | KEITHLEY SOURCEMETER | DR VOON | INEE 0000095-250414
96 | LAI SIANG LENG | FYP |  | MICRO E | EVAPORATOR ONLY | ROSYDI/DR FOO | INEE 0000096-250414
97 | LAU KUEN YAU | FYP |  | PHOTONIC ENGINEERING | ONLY 3D PROFILER | DR MUKHZER/? | INEE 0000097-060514
98 | MOHD RIDZUAN BIN MUDA | MSc |  | MICRO E | CENTRIFUGE/EDX/SEM | MUHAMMAD MAHYUDDIN RAMLI | INEE 0000098-060514
99 | SUHAIMI BIN MAT KAHAR | FYP | 016-3052747 | MICRO E | ONLY EVAPORATOR | DR YUFRIDIN | INEE 0000099-220514
100 | NURHAZWANI BINTI MUSA | MSc | 013-792-7508 | MATERIAL | LT FUR/SPIN COAT/PVD EVA | PROF DR MOHD NOOR AHMAD | INEE 0000100-020714
101 | NORHAYATI BINTI MOHD NASIR | MSc | 013-9763262 | MATERIAL | ONLY 4 POINT PROBE | DR NORAINIZA SAUD | INEE 0000101-090714
102 | NURUL'ATIRAH BINTI MD NOOR | MSc | 016-480-8580 | MATERIAL | Only potentiostat | SITI NORBAHIYAH/DR NAZREE | INEE 0000102-140714
103 | NAZIRA BINTI MAT AMIN | MSc | 013-412-3756 | MATERIAL | Only potentiostat | DR NAZREE | INEE 0000103-260814
104 | ABDUL KADER ALAKRACH | MSc | 013-942-4050 | POLYMER | SAMPLES FOR SEM | DR AZLIN FAZLINA OSMAN | INEE 0000104-270814
105 | GOMESH NAIR A/L SHASIDHARAN | PhD | 012-256-2015 | SYSTEM E | SEM/DIELECTRIC ANALYZER/3D PROFILER | DR. MUHAMMAD IRWANTO/PROF UDA | INEE-0000105-050914
106 | ROSHILA BINTI MOHD LAZIM | FYP | 013-417-4681 | MICRO E | ALL | PROF UDA | INEE-0000106-131014
107 | CHE NORIANI BT HUSSAIN | FYP | 014-292-8084 | MICRO E | ALL | PROF UDA | INEE-0000107-131014
108 | SITI RAQEEMA BINTI ALI | FYP | 014-599-0961 | MICRO E | ALL | PROF UDA | INEE-0000108-131014
109 | AHMAD LABIB BIN CHE AZIZ | FYP | 011-1630-8676 | MICRO E | HOT PLATE/SPIN COATER ONLY | HARIYANTI/DR FOO | INEE-0000109-231014
110 | MUHSON NAJIY BIN ROSHIDEE | FYP | 013-402-3513 | MICRO E | HOT PLATE/SPIN COATER ONLY | HARIYANTI/DR FOO | INEE-0000110-231014
111 | ADELYN PUAH YING PHING | PhD | 012-489-1829 | INEE | ALL | PROF UDA | INEE-0000111-231014
112 | TONY YOO CHUNG SUNG | MSc | 014-398-4186 | INEE | ALL | DR VOON | INEE-0000112-281014
113 | SUHAIMI BIN MAT KAHAR | MSc | 016-305-2747 | INEE | ALL | DR VOON | INEE-0000113-281014
114 | HA YONG PENG | FYP | 016-543-5793 | MICRO E | ALL | PROF UDA | INEE-0000114-041114
115 | MOHD IZRUL IZWAN B RAMLI | MSc | 013-278-6509 | MATERIAL | ONLY 4 POINT PROBE | DR NORAINIZA SAUD/DR NAZREE | INEE-0000115-041114
116 | THEY YEE CHIN | MSc | 019-448-6628 | MICRO E | OXIDE FUR/ANNEAL/SEM/METAL DEP | PM DR ZALIMAN | INEE-0000116-101114
117 | SIN LYE LENG | FYP | 014-603-5835 | MICRO E | ALL | DR KHAIRUDDIN | INEE-0000117-121114
118 | TEOH CHAI LING | FYP | 016-415-8489 | MICRO E | ALL | DR KHAIRUDDIN | INEE-0000118-121114
119 | PENG WEI KEE | FYP | 017-486-4726 | MICRO E | ALL | DR FOO KAI LOONG | INEE-0000119-191114
120 | LEONG CHIE HOU | FYP | 012-616-1516 | MICRO E | ALL | DR FOO KAI LOONG | INEE-0000120-191114
121 | SOO HOO KEAN AUN | FYP | 017-412-8660 | MICRO E | ALL | DR FOO KAI LOONG | INEE-0000121-191114
122 | CHENG QING ZHOU | FYP | 016-887-7106 | MICRO E | ALL | DR FOO KAI LOONG | INEE-0000122-191114
123 | NIK KHAIRUN SYAHIRAH BT NORDIN | MSc | 013-586-4752 | INEE | ALL | PROF UDA | INEE-0000123-191114
124 | THIVINA VIJAYAKUMARAN | MSc | 016-406-6586 | INEE | ALL | PROF UDA | INEE-0000124-191114
125 | SITI NURDIYANA BT ABD RASHID (tweezer) | FYP | 019-567-1040 | MICRO E | ALL | DR RUSLINDA | INEE-0000125-191114
126 | FADHLINA BINTI MOKHTAR (tweezer) | FYP | 017-422-1783 | MICRO E | ALL | DR RUSLINDA | INEE-0000126-191114
127 | CHANG HUI YI | FYP | 017-286-9823 | MICRO E | ALL | DR KHAIRUDDIN | INEE-0000127-241114
128 | MUHAMMAD NAZRIN ASYRAF BIN OMAR | FYP | 017-526-5350 | MICRO E | Spinner/Hot plate/Mask Aligner | DR SHAZLINA JOHARI/NUZAIHAN | INEE-0000128-241114
129 | FATIN NABILAH BINTI MOHD HASAN | FYP | 017-290-4864 | MICRO E | ALL | DR FOO KAI LOONG | INEE-0000129-281114
130 | FATINHANANI BINTI ABD RAHMAN | FYP | 017-557-9919 | MICRO E | ALL | DR FOO KAI LOONG | INEE-0000130-281114
131 | NURISMALIZA BINTI JURAMI (tweezer) | FYP | 019-405-5097 | MICRO E | ALL | DR RUSLINDA | INEE-0000131-281114
132 | MOHD HAIRI BIN MAHAYUDIN | FYP | 014-245-4724 | MICRO E | ALL | PROF UDA | INEE-0000132-151214
133 | Dr Abadal Salam Taha Hussain | Reseacher | 011-1109-9077 | Electrical System | All | NA | INEE-0000133-190115
134 | NURUL FAZLINDA BT MOHAMED RANI | FYP | 014-522-5812 | MICRO E | All | RAMZAN | INEE-0000134-210115
135 | AHMAD HELMI BIN MAHAMAROWI | FYP | 017-672-0634 | MICRO E | All | RAMZAN | INEE-0000135-210115
136 | Wrong key-in. | Wrong key-in. | Wrong key-in. | Wrong key-in. | Wrong key-in. | Wrong key-in. | Wrong key-in.
137 | SOH ING TIAN | FYP | 017-531-8583 | MICRO E | Only LPCVD | PM YUFRIDIN | INEE-0000137-270115
138 | AINUR FATINI BINTI SHAIDIN | FYP | 012-591-1625 | MATERIAL/POLYMER | Only I-V | DR IRWANA NAINGGOLAN | INEE-0000138-270115
139 | MOHD LOKMAN HAKIM BIN GHAZALY | FYP | 013-960-3161 | MATERIAL/POLYMER | Only I-V | DR IRWANA NAINGGOLAN | INEE-0000139-270115
140 | AIMAN HAZIQ BIN AZADDIN | FYP | 013-798-0830 | MICRO E | All | RAMZAN | INEE-0000140-280115
141 | NORLIDA BINTI ABU BAKAR | MSc | 017-523-7877 | MICRO E | Only Evaporator | PM DR JOHARI | INEE-0000141-300115
142 | MUHAMMAD FATHI BIN MD TAIB | FYP | 014-254-4123 | MICRO E | All | RAMZAN | INEE-0000142-300115
143 | GARY BIN ASIP | FYP | 014-943-8992 | MICRO E | All | RAMZAN | INEE-0000143-300115
144 | TENGKU NUR ATIQAH BT TENGKU AB MUTALIB | MSc | 013-613-6374 | MATERIAL | Only I-V & Yellow booth | Dr TAN SOO JIN/DR FOO | INEE-0000144-030315
145 | NORHANISAH BT ZAKARIA | FYP | 011-3308-7223 | MICRO E | Only Vacuum Oven & UV-VIS | DR MAHYIDDIN | INEE-0000145-031915
146 | NOR IMAN BIN ZAABA | MSc | 013-407-2031 | MICRO E | ALL | DR FOO KAI LOONG | INEE-0000146-032515
147
148
149
150
151
152
153
154
155
156
157
Note on C122: ismail - [2010]: borrow labcoat
Note on C123: ismail - [2010]: borrow labcoat
Note on C124: ismail - [2010]: borrow labcoat
Note on C125: ismail - [2010]: borrow labcoat
Note on C149: User: Borrow labcoat

## Budget Maintenance Equipment
MACHINE/EQUIPMENT | PARTS OF MAINTENANCE | SUPPLIER
LPCVD | 1- PARTS OF MACHINE | JOHARI HITECH Craft (M) Sdn Bhd
2- WIRING 3 PHASE
3- THERMAL COUPLE
4- GAS N2 PURGE/VENT
THERMAL EVAPORATOR | 1- SHIELD X 3 | SHARUL ULVAC
2-Turbo pump swap
3-LPCVD heater 3 phase wiring
4-LPCVD heater thermocouple calibration
5-Evaporator shuttle motor
RTA | 1- | YAZID NANORIAN
2-
ICP-RIE | 1- FINGER PLATE | YAZID NANORIAN
2- CERAMIC COVER
OXIDATION FURNACE | 1- GLASS BOAT 4 INCH / 2 INCH | JOHARI HITECH Craft (M) Sdn Bhd
MASK ALIGNER | 1-UV LIGHT | YAZID NANORIAN
2- SERVICE SPIN COATER

## CHEMICAL REGISTER
LAST UPDATE 25032013-REFER CONSIGNMENT SYSTEM
CHEMICAL REGISTER LIST
LOCATION: UNIMAP NANO BIOCHIP LABORATORY.
WORK UNIT: NANO BIOCHIP MATERIAL
No of chemical | Product name | Chemical Composition | CAS No | Maximum Stock | Status Chemical | Source of Information | Physical chemical | Hazard category | Skin Notation (Y/N) | Label (Y/N) | Name, address of supplier and contact number.(Tel. No E-mail) | contact vandor | Owner | Storage/Rack
1 | Aluminium Nitrate Nonahydrate | Aluminium Nitrate Nonahydrate | 7784-27-2 | 10G 1 QTY | Active | MSDS | Powder | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
2 | Palladium (ii) Chloride | Palladium Chloride | 7647-10-1 | 1G 2QTY | Active | MSDS | Powder | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
3 | Cobalt (ii) Chloride Hexahydrate | Cobalt (ii) Chloride Hexahydrate | 7791-13-1 | 25G 1QTY | Active | MSDS | Powder | Extremely Flammable | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
4 | Acetonitrile | Acetonitrile | 75-05-8 | 100ML 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
5 | Zinc Chloride | Zinc Chloride | 7646-85-7 | 250G 1QTY | Slow usage | MSDS | Powder | Harmful | Y | Y | Vandor Name : Che Umar Bin CheLah Add : 311 F, Tingkat 1, Bangunan Zakat, Jalan Kampung Baru, 08000 Sungai Petani, Kedah Darul Aman | Telephone : 019-5907788 | Kashif | Yellow
6 | Zinc Acetate Dihydrate | Zinc Acetate Dihydrate | 5976-45-6 | 250G 1QTY | Active | MSDS | Powder | Harmful | Y | Y | Vandor Name : Che Umar Bin CheLah Add : 311 F, Tingkat 1, Bangunan Zakat, Jalan Kampung Baru, 08000 Sungai Petani, Kedah Darul Aman | Telephone : 019-5907789 | Kashif | Yellow
7 | Nickel (ii) Acetate Tetrahydrate, 99.998% Metal Basic | Nickel di(acetate) | 6018-89-9 | 10G 1 QTY | Active | MSDS | Powder | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
8 | Silvernitrate, › 99.9% | Silvernitrate | 7761-88-8 | 25G 1QTY | Active | MSDS | Powder | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
9 | Zincnitrate Hexahydrate, Reagent Grade 98% | Zincnitrate Hexahydrate | 10196-18-6 | 500G 2QTY | Active | MSDS | Powder | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
10 | Gold (iii) Chloride Hydrate | chloroauric acid, hydrogen tetrachloroaurate (iii) hydrate, tetachlovoaurics(iii) acid | 27988-77-8 | 5G 1QTY | Active | MSDS | Powder | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
11 | Imidazole | 1,3-diaza-2, 4 cyclopentadience, glyoxaline | 288-32-4 | 5G 1QTY | Slow usage | MSDS | Powder | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Yacub | Yellow
12 | 6-Mercapto-1- hexanol 97% | 6-Mercapto-1- hexanol | 1633-78-9 | 5ML 1QTY | Slow usage | MSDS | Liquid | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Yacub | Yellow
13 | Hexamethylenetetramine,Reagent Plus 99% | Urotropne, 1,3,5,7-Tetraazatricyclo, Hexamine, Methenamine | 100-97-0 | 500G 1QTY | Active | MSDS | Powder | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
14 | Potassium Chloride | Potassium Chloride | 7447-40-7 | 1KG 2QTY | Active | MSDS | Powder | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
15 | Iron powder, -325 meshm, 97% | Iron Powder | 7439-89-6 | 500G 1QTY | Slow usage | MSDS | Powder | Flammable | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
16 | Potassium Hydroxide | Potassium Hydroxide | 1310-58-3 | 2QTY | Active | MSDS | Powder | Flammable | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353322 | Kashif | Yellow
17 | Diethylamine | Diethylamine | 109-89-7 | 1L 1QTY | Slow usage | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
18 | Diethanolamine | Bis(2-hydroxyethyl)amine 2,2′-Iminodiethanol | 111-42-4 | 1L 1QTY | Slow usage | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
19 | 2-Propanol, anyhydrous 99.5% | sec-Propyl alcohol Isopropyl alcohol Isopropanol | 67-63-0 | 1L 1QTY | Active | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
20 | 2-Methoxytehanol | Methyl Cellosolve Methyl glycol Ethylene glycol monomethyl ether | 109-86-4 | 1L 1QTY | Active | MSDS | Liquid | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
21 | Tin (ii) Chloride | Stannous chloride | 7772-99-8 | 100 G 1QTY | Active | MSDS | Powder | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
22 | a-Terpineol | (S)-2-(4-Methyl-3-cyclohexenyl)-2-propanol (S)-p-Menth-1-en-8-ol | 10482-56-1 | 50 ML 1 QTY | Slow usage | MSDS | Liquid | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Kashif | Yellow
23 | Ethanolamine | 2-Aminoethanol, Monoethanolamine | 141-43-5 | 1L 2QTY | Slow usage | MSDS | Liquid | Corrosive | Y | Y | Vandor Name : Che Umar Bin CheLah Add : 311 F, Tingkat 1, Bangunan Zakat, Jalan Kampung Baru, 08000 Sungai Petani, Kedah Darul Aman | Telephone : 019-5907788 | Kashif | Yellow
24 | Titanium (iv) isopropoxide | Tetraisopropyl orthotitanate | 546-68-9 | 500ML 1QTY | Active | MSDS | Liquid | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Sharifah | Yellow
25 | Tris (2-Aminoethyl) Amine 96% | Tris (2-Aminoethyl) Amine | 4097-89-6 | 100ML 1QTY | Active | MSDS | Liquid | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353322 | Sharifah | Yellow
26 | Titanium (iv) butoxide | Orthotitanic acid tetrabutylester Tetrabutyl orthotitanate Tetrabutyl titanate TYZOR® TBT organic titanate TNBT | 5593-70-4 | 1L 1QTY | Active | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Sharifah | Yellow
27 | phosphate buffer solution in water pH 7.2 | Sodium carbonat | 497-19-8 | 2L 1QTY | Slow usage | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Foo Kai Long | Yellow
28 | Zinc | zinc | 7440-66-6 | 5G 1QTY | Slow usage | MSDS | powder | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Foo Kai Long | Yellow
29 | Sodium Citrate | Trisodium citrate dihydrate Citric acid trisodium salt dihydrate | 6132-04-3 | 1KG 1QTY | Slow usage | MSDS | powder | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Foo Kai Long | Yellow
30 | Glycine | Aminoethanoic acid Aminoacetic acid Glycocoll | 56-40-6 | 100G 1QTY | Active | MSDS | powder | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Weng | Yellow
31 | Sodium cyanoborohydrade | Sodium cyanoborohydrade | 25895-60-7 | 100G 1QTY | Active | MSDS | Powder | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Weng | Yellow
32 | Sodium hydroxide | Caustic soda | 1310-73-2 | 100G 1QTY | Active | MSDS | Powder | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Weng | Yellow
33 | Glutaraldehydate solution grade ii, 25% in H2O | Glutaric dialdehyde solution Gluteraldehyde Pentane-1,5-dial | 111-30-8 | 100ML 1QTY | Slow usage | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Weng | Yellow
34 | Glutaraldehydate solution | Glutaric dialdehyde solution Gluteraldehyde Pentane-1,5-dial | 111-30-8 | 250ML 1QTY | Slow usage | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Weng | Yellow
35 | Phosphate Buffer Powder, 0.1M | Phosphate Buffer Powder | N/A | 3.8ML 1QTY | Slow usage | MSDS | Liquid | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Weng | Yellow
36 | 3- Aminopropyl-Trimethoxysilan, 97% | 3- Aminopropyl-Trimethoxysilan | 13822-56-5 | 100 ML 1QTY | Slow usage | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Weng | Yellow
37 | Toluene, Anhydrous, 99.8% | Toluene, Anhydrous | 108-88-3 | 1L 1QTY | Slow usage | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353321 | Weng | Yellow
38 | Glutaraldehyde Solution | Glutaric dialdehyde solution Gluteraldehyde Pentane-1,5-dial | 111-30-8 | 1L 1QTY | Slow usage | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353322 | Weng | Yellow
39 | Sodium Hydroxide Pellet | Caustic soda | 1310-73-2 | 1KG 1QTY | Slow usage | MSDS | Pellet | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353323 | Weng | Yellow
40 | mr- Rem 660 Remover for photoresist | N-methyl-2-pyrrolidone | 872-50-4 | 5 L 1QTY | Active | MSDS | Liquid | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353324 | Foo Kai Long | Blue
41 | ma-P 1200 Positive Tone Photoresist | Anisole, 2-Methoxy-1-Methylethyl acetate | 108-65-6 | 0.5 L 1QTY | Active | MSDS | Liquid | Irritant | y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353325 | Foo Kai Long | Blue
42 | ma-N 1400 Negative Tone Photoresist | Anisole, 2-Methoxy-1-Methylethyl acetate | 100-66-3, 108-65-6 | 1.0L 1QTY | Active | MSDS | Liquid | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353326 | Foo Kai Long | Blue
43 | Organic Polymer Solution, | Cyclopentanone | 120-92-3 | 1x500ml 1QTY | Slow usage | MSDS | Liquid | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353327 | Foo Kai Long | Blue
44 | EBR PG | 1,3-Dioxolane, Propylene glycol monomethyl | 6646-06-0, 107-98-2 | 1X4L 1QTY | Slow usage | MSDS | Liquid | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353328 | Foo Kai Long | Blue
45 | PMGI 101A Developer | Tetraethyammonium Hydroxide, water | 77-98-5, 7732-18-5 | 1X4L 1QTY | Slow usage | MSDS | Liquid | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353329 | Foo Kai Long | Blue
46 | Ma-D 5335 Developer for Photoresist | Tetramethlammonium Hydroxide | 75-59-2 | 5L 1QTY | Slow usage | MSDS | Liquid | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353330 | Foo Kai Long | Blue
47 | Remover PG Photoresist Remover |  | N/A | 4L 1QTY | Slow usage | MSDS | Liquid | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353331 | Foo Kai Long | Blue
48 | ma-D 381 Developer for photoresist | Caustic soda, sodium hydroxide | 215-185-5 | 5L 1QTY | Slow usage | MSDS | Liquid | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353332 | Foo Kai Long | Blue
49 | Ethanol 95% |  | N/A | 2.5L 1QTY | Active | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353333 | Foo Kai Long | Blue
50 | Phosphorus Pentoxide | Phosphoric anhydride Phosphorus(V) oxide | 1314-56-3 | 5G 1QTY | Active | MSDS | Powder | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353322 | Kashif | Yellow
51 | Acetone | Acetone | 67-64-1 | 1 GALL 5 QTY | Active | MSDS | Liquid | Extremely Flammable | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353323 | INEE | Cabimet (A)
52 | Acetone | Acetone | 67-64-1 | 2.5 L 7 QTY | Active | MSDS | Liquid | Extremely Flammable | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353324 | INEE | Cabimet (A)
53 | 2-Propanol | 1-Methoxy-2-propanol, 2-Methoxy-1-propanol | 107-98-2, 1589-47-5 | 1 GALL 7QTY | Active | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353325 | INEE | Cabimet (A)
54 | 2-Propanol | 1-Methoxy-2-propanol, 2-Methoxy-1-propanol | 107-98-2, 1589-47-6 | 2.5 L 1QTY | Active | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353326 | INEE | Cabimet (A)
55 | Methanol | Methanol | 67-56-1 | 2.5 L 1QTY | Active | MSDS | Liquid | Irritant Flammable | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353327 | INEE | Cabimet (A)
56 | Ethanol 95 % | Ethyl Alcohol, Isopropyl Alcohol, Methyl Alcohol | 64-17-5, 67-63-0 | 2.5 L 1 QTY | Active | MSDS | Liquid | highly flammable | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353328 | INEE | Cabimet (A)
57 | Aluminium Etch | Nitric Acid , Acetic Acid, Phosphoric Acid, Water | 7697-37-2, 64-19-7, 7664-38-2, 7732-18-5 | 2.5 L 3QTY | Active | MSDS | Liquid | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353329 | INEE | Cabimet (A)
58 | Hydrochloric acid fuming 37% | Hydrochloric acid | 7647-01-0 | 2.5 L 1 QTY | Active | MSDS | Liquid | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353330 | INEE | Cabimet (A)
59 | Analytical reagent grade Toluene | Toluene | 108-88-3 | 2.5 L 1 QTY | Slow usage | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353331 | INEE | Cabimet (A)
60 | Hydrogen Peroxide 30% | Hydrogen Peroxide, water | 7722-84-1, 7732-18-5 | 2.5 L 1 QTY | Slow usage | MSDS | Liquid | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353332 | INEE | Cabimet (A)
61 | Sulfuric Acid 95-97% | sulfuric acid | 7664-93-9 | 2.5 L 1 QTY | Slow usage | MSDS | Liquid | Corrosive | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353333 | INEE | Cabimet (A)
62 | Buffer solution pH 1.00 | HYDROCHLORIC ACID | 7647-01-0 | 1.0L 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353334 | INEE | FREEZER
63 | Buffer solution pH 2.00 | HYDROCHLORIC ACID | 7647-01-0 | 1.0L 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353335 | INEE | FREEZER
64 | Buffer solution pH 3.00 | HYDROCHLORIC ACID | 7647-01-0 | 1.0L 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353336 | INEE | FREEZER
65 | Buffer solution pH 4.00 | HYDROCHLORIC ACID | 7647-01-0 | 1.0L 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353337 | INEE | FREEZER
66 | Buffer solution pH 5.00 | HYDROCHLORIC ACID | 7647-01-0 | 1.0L 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353338 | INEE | FREEZER
67 | Buffer solution pH 6.00 | HYDROCHLORIC ACID | 7647-01-0 | 1.0L 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353339 | INEE | FREEZER
68 | Buffer solution pH 7.00 | HYDROCHLORIC ACID | 7647-01-0 | 3X500 ML (2QTY) | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353340 | INEE | FREEZER
69 | Buffer solution pH 8.00 | HYDROCHLORIC ACID | 7647-01-0 | 1.0L 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353339 | INEE | FREEZER
70 | Buffer solution pH 9.00 | HYDROCHLORIC ACID | 7647-01-0 | 1.0L 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353340 | INEE | FREEZER
71 | Buffer solution pH 10.00 | HYDROCHLORIC ACID | 7647-01-0 | 1.0L 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353341 | INEE | FREEZER
72 | Buffer solution pH 11.00 | HYDROCHLORIC ACID | 7647-01-0 | 1.0L 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353342 | INEE | FREEZER
73 | Buffer solution pH 12.00 | HYDROCHLORIC ACID | 7647-01-0 | 1.0L 1QTY | Active | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353343 | INEE | FREEZER
74 | HI 7010 pH 10.01 Buffer Solution ± 0.01 ph | Aqueous buffer solution | N/A | 1 L 1 QTY | Slow usage | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353344 | INEE | FREEZER
75 | HI 7004 pH 4.01 Buffer Solution ± 0.01 ph | Aqueous buffer solution | N/A | 1 L 1 QTY | Slow usage | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353345 | INEE | FREEZER
76 | HI 7007 pH 7.01 Buffer Solution ± 0.01 ph | Phosphate compounds: <1% Other ingredients: stabilizers: <0.01% | N/A | 1 L 1 QTY | Slow usage | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353346 | INEE | FREEZER
77 | HI 7007 pH 7.01 Buffer Solution ± 0.01 ph | Phosphate compounds: <1% Other ingredients: stabilizers: <0.01% | N/A | 500ml 1 QTY | Slow usage | MSDS | Liquid | Toxic | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353347 | INEE | FREEZER
78 | sodium sulfate-10-hydarte | Sodium sulfate | 7727-73-3 | 1KG 3QTY | Slow usage | MSDS | Powder | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353348 | Dr Nazree | Cabimet (B)
79 | Aluminium Powder | Aluminium | 7429-90-5 | 500G 1QTY | Slow usage | MSDS | Powder | Harmful | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353349 | Dr Nazree | Cabimet (B)
80 | Oxalic Acid-2-Hydrate | Oxalic acid | 6153-56-6 | 500G 5QTY | Slow usage | MSDS | Powder | Irritant | Y | Y | Add: Sigma-Aldrich (M) Sdn. Bhd, A-07-11, Empire Office, Empire Subang, Jalan SS16/1, SS16, 47500 Subang Jaya- Selangor Darul Ehsan Malaysia. | Telephone : +60 (603)56353350 | Dr Nazree | Cabimet (B)